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THE UNIVERSITY OF NEW MEXICO ¢ HEALTH SCIENCES CENTER

UNM HOSPITALS




 




PURCHASING DEPARTMENT
2211 Lomas Blvd., NE

Albuquerque, New Mexico 87106
REQUEST FOR PROPOSAL NUMBER: 140-09   DUE DATE: 6/16/09
TIME: 3:00 PM MST
TITLE:   A Bi-plane Interventional Room Solution for Interventional Radiology 

University of New Mexico Hospitals.  (UNMH) invites you to submit a proposal for A Bi-plane Interventional room solution for Interventional Radiology.  Please read carefully all instructions, specifications, terms and conditions. Failure to comply with the instructions, terms and conditions, and specifications of this Request for Proposal may result in your proposal being classified as non-responsive.  Proposals must be submitted to the above address no later than the time and date specified above.  New Mexico criminal law prohibits bribes, gratuities and kickbacks.

PROPOSAL SUBMITTAL

All information must be entered in ink or typed.  Corrections must be initialed by individual signing the offer.  The proposal title, number, opening date and time must be clearly indicated in the lower left corner of your sealed proposal envelope.  Failure to properly label your envelope may result in your proposal not being included in the proposal opening and evaluation.
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CONTACT INFORMATION:  All questions must be submitted to both contacts via email
BUYER:          
 Lynn Nelson    (505) 272-9156
Email: Salnelson@salud.unm.edu
TECHNICAL:  
 Gordon Weimer 


Email: Gweimer@salud.unm.edu

OTHER INFORMATION:  Deadline for submission of questions, Friday, May 22  4:00 PM
NUMBER OF COPIES REQUIRED: ​​3- Hard Copies and 1-CD
SCHEDULED SITE VISITS: TBA
FREIGHT TERMS: Destination
PAYMENTS TERMS: Net 30 Days

TERMS AND CONDITIONS

The purchase of materials and/or services awarded under this Request for Proposal are subject to the University of New Mexico's General Terms and Conditions, The Standard Terms and Conditions for goods and services, in addition to other statements contained in this Request for Proposal.

UNM Hospitals
Purchasing Department
2211 Lomas Blvd., NE

Albuquerque, New Mexico 87106

REQUEST FOR PROPOSAL NUMBER: 140-09       DUE DATE: 6/16/09
TIME: 3:00 PM MST
TITLE:   A Bi-plane Interventional Room Solution for Interventional Radiology
The Following Information Must Be Completed And Returned With Your Proposal:
TELEPHONE NUMBER FOR OFFER CLARIFICATION:

NAME: ___________________________________ 
EMAIL: _________________________________

TITLE: ___________________________________        FAX NUMBER: ___________________________
ACKNOWLEDGMENT OF ADDENDA:

The undersigned acknowledges receipt of the following addenda:

Addenda No.  ________ Dated:  ________

 Addenda No.  _______ Dated: _______

Addenda No.  ________ Dated:  ________ 
 Addenda No.  _______ Dated: _______

AGREEMENT TO SUPPLY BONDS:

The undersigned agrees to furnish Bonds, if requested, by this Request for Proposal at no cost to UNM Hospitals.

AUTHORIZED REPRESENTATIVE:

The undersigned, as an authorized representative for the Company named below, acknowledges that he/she has examined this Request for Proposal with its related documents, and being familiar with all of the conditions surrounding the described materials and/or services, including the availability of materials and labor, hereby offers to furnish all labor, materials, and supplies necessary to comply with the requirements and specifications in accordance with the Terms and Conditions set forth herein and at the prices stated in the price quote.

COMPANY NAME: __________________________________________________________________

ADDRESS: _________________________________________________________________________

CITY/STATE/ZIP: __________________________________________________________________


SIGNATURE OF AUTHORIZED REPRESENTATIVE: ____________________________________
PRINTED OR TYPED NAME: _________________________________________________________
TITLE: _____________________________________________________DATE: _________________
 PHONE:  (____) ___________ FAX:  (_____) __________EMAIL: _____________________________

CONTRACTOR’S LICENSE NUMBER: _________________________________________________
CONTRACTOR’S FEDERAL TAX ID NUMBER__________________________________________

NEW MEXICO GROSS RECEIPTS TAX NUMBER________________________________________
SECTION I

UNIVERSITY OF NEW MEXICO - UNM HOSPITALS

GENERAL TERMS AND CONDITIONS
1.
ADDRESSES FOR NOTICES: Any notice required to be given under this proposal or the resultant contract shall be in writing and delivered in person or via first class mail or overnight mail to the following address:

Mail to:





Deliver to:
UNM Hospitals




UNM Hospitals
Purchasing Department



Purchasing Department

2211 Lomas Blvd. NE



2211 Lomas Blvd. NE

Albuquerque, NM 87106



Albuquerque, NM  87106

2.
AGREEMENT TERMS AND CONDITIONS: The agreement between the University of New Mexico and a Seller will follow the format specified by the University of New Mexico and contain the terms and conditions set forth in “Terms and Conditions”. However, the University of New Mexico reserves the right to negotiate with a successful Seller provisions in addition to those contained in this RFP. The contents of this RFP, as revised and/or supplemented, and the successful seller’s proposal will be incorporated into the contract.

 Should a Seller object to any of the University’s terms and conditions, as contained in this Section, that Seller must propose specific alternative language that would be acceptable to the University. General references to the Sellers terms and conditions or attempts at complete substitutions are not acceptable to the University of New Mexico and will result in disqualification of the Seller’s proposal.  Sellers must provide a brief discussion of the purpose and impact, if any, of each proposed change followed by specific proposed alternate wording.

3.
APPROPRIATION:  The terms of the contract are contingent upon sufficient appropriations and authorization being made by the Regents of the University of New Mexico.  If sufficient appropriations and authorization are not made by the University, the contract shall, not withstanding any other provisions of the contract, terminate immediately upon the Contractor’s receipt of written notice of termination from the University.

4.
CANCELLATION:  UNM Hospitals reserves the right to cancel without penalty, this request, the resultant contract or any portion thereof for unsatisfactory performance, cancellation of the project or unavailability of funds.

5.
CHANGES/ALTERATIONS AFTER AWARD: Changes or alterations after the award can only be made if agreed to in writing by the University.

6.
CLEAN UP: It is the Seller’s responsibility that the job site be kept clean and free of rubble while work is performed under this contract. Upon completion of the work, all areas shall be cleared of all contractor’s equipment, excess materials and rubble.

7.
CONFLICT OF INTEREST FORM: Seller is required to sign the attached Supplier Conflict Of Interest and Debarment and Suspension Certification FORM. Failure to provide the University of New Mexico (UNM) with a completed Conflict of Interest Form may result in your offer being considered non-responsive.

8.
CONSTRUCTION SCHEMATIC DRAWINGS: The successful Seller will be required to provide the University, upon completion of the work, three (3) sets of schematic drawing(s) showing electrical, plumbing, utility lines, etc. added or effected by the work the contractor has performed, unless other provisions for preapproval of drawings is contained elsewhere in this Proposal.

9.
COOPERATION AND DISPUTE RESOLUTION: The parties agree that, to the extent compatible with the separate and independent management of each, they will maintain effective liaison and close cooperation.  If a dispute arises related to the obligations or performance of either party under this Agreement, representatives of the parties will meet in good faith to resolve the dispute.

10.
COST ANALYSIS/BREAKDOWN REQUIRED: A cost analysis or breakdown of the offered cost is required to be submitted with your response. The cost analysis information may be marked as confidential information if it includes information you feel is proprietary to your operation. The total offered cost cannot be considered confidential.

11.
DAMAGE AND SECURITY OF UNM PROPERTY: The Seller shall be responsible for all damage to persons or property that occurs as a result of his fault or negligence, or that of any of his employees, agents and/or subcontractors. He shall save and keep harmless the University against any and all loss, cost, damage, claims, expense or liability in connection with the performance of this contract. Any equipment or facilities damaged by the Seller’s operations shall be repaired and/or restored to their original condition at the Seller’s expense, including but not limited to cleaning and painting. The Seller shall be responsible for security of all his equipment and for the protection of work done under this contract until final acceptance of the work.

12
DELIVERY DATE: Delivery is an important consideration and is a factor in determining the award. If you cannot meet the delivery date stated on the cover sheet, please state your earliest delivery date in your offer.

13.
DISCLOSURE OF PROPOSAL CONTENTS: The proposals will be kept confidential until the University of New Mexico (UNM) awards a price agreement. At that time, all proposals and documents pertaining to the proposals will be open to the public, except for the material that is proprietary or confidential. The Procurement Managers will not disclose or make public any pages of a proposal on which the Seller has stamped or imprinted “proprietary” or “confidential” subject to the following requirements. 

14.
Proprietary or confidential data shall be readily separable from the proposal in order to facilitate eventual public inspection of the non-confidential portion of the proposal. Confidential data is normally restricted to confidential financial information concerning the Seller’s organization and data that qualifies as a trade secret in accordance with the Uniform Trade Secrets Act, 57-3A-l to 57-3A-7 NMSA 1978. The price of products offered or the cost of services proposed shall not be designated a proprietary or confidential information. 

If a request is received for disclosure of data for which a Seller has made a written request for confidentiality, the University of New Mexico (UNM) shall examine the Seller’s request and make a written determination that specifies which portions of the proposal should be disclosed. Unless the Seller takes legal action to prevent the disclosure, the proposal will be so disclosed. The proposal shall be open to public inspection subject to any continuing prohibition on the disclosure of confidential data.  
15.
DISRUPTION OF NORMAL ACTIVITY: All work shall be performed so as not to interfere with normal   University activities. When it is necessary to disrupt normal activities, the schedule of work and the areas to be affected must be approved by the University’s authorized representative prior to commencement of the work.

16.
 EMPLOYEE CERTIFICATION: The Seller and all Sellers’ employees utilized on the work to be performed under this Request must have the proper certification(s) and license(s) to comply with State and local requirements connected to this Request. The Seller shall use only fully qualified and approved service technicians to perform inspections, service and/or repairs under this request.  

16.
 EQUIPMENT REQUIRED: The Seller shall be responsible for supplying and maintaining all equipment and materials necessary to complete the work under this Request except as otherwise noted in the Specifications.

17.
 FINANCIAL STATEMENT REQUIRED: All offerors are required to submit with their offer a Financial Statement from the end of their most recent fiscal year.

18.
GENERAL TERMS AND CONDITIONS: The University of New Mexico’s General Terms and Conditions are an equal and integral part of this request.  All terms and conditions of this request will remain unchanged for the duration of the contract and will supersede and take precedence over any vendor agreement forms.  Offeror must include a detailed description regarding any exceptions to the terms and conditions of this request for proposal.  If exceptions or deviations are not clearly stated, it is understood that the terms and conditions of this proposal shall govern.   The University of New Mexico reserves the right to reject any proposal that does not meet the terms and conditions of the request for proposal.  It further reserves the right to accept any modifications to the terms and conditions if it is in the best interest of the University to do so.

10.
GOVERNING LAW: This proposal and the resultant contract and/or purchase order will be interpreted and governed by the laws of the State of New Mexico. 
20.
INSURANCE REQUIRMENTS: The Seller is required to carry insurance meeting the requirements in the Exhibit labeled “INSURANCE REQUIREMENTS” or as noted in the specifications. Seller must submit the Certificate of Insurance to the appropriate Buyer prior to commencing work under this contract. Insurance shall remain in effect for the entire term of the contract and must be extended to coincide with any future contract extensions. This Request for Proposal Number must appear on the Certificate of Insurance.

21.
MEMBERSHIP IN PURCHASING GROUPS: UNM Hospitals is eligible and will at its discretion, accept Novation, GSA SPD, or UHC contract pricing, where applicable, and where it is the lowest price.

22.
MERGER:  The contract shall incorporate all the agreements, covenants, and understandings between the parties thereto concerning the subject matter thereof.  No prior agreements or understandings, verbal or otherwise, of the parties or their agents shall be valid or enforceable unless embodied in the contract.

23.
MULTI-TERM CONTRACT: Pursuant to Section 13-1-151 NMSA 1978, it has been determined that the estimated requirements covered in this contract are reasonably firm and continuing; and that a multi-term award will serve the best interests of the University of New Mexico (UNM).  

24.
NEW MATERIALS REQUIRED: All materials and equipment delivered and/or installed under this Request shall be new and be the standard products of a manufacturer regularly engaged in the production of the materials and equipment. Where two or more units of the same class of materials and/or equipment are required, the units shall be the products of the same manufacturer. Any manufacturer’s data supplied with the item(s) shall be submitted to the University’s authorized representative.  
25.
NON-PERFORMANCE PENALTIES: The Seller agrees to pay the University an amount equal to $50.00 per day for each calendar day past the completion date specified in this contract that completion or delivery is delayed. The University may subtract this amount from any monies due to the Seller.

26.
OPTION TO RENEW: The University reserves the option to renew the resultant contract if such renewal is mutually agreed to and found to be in the best interests of the University. These renewal options will be exercised in increments as indicated in the proposal specifications, or if not stated, in one-year terms. The contract shall not exceed (4) four years including all renewals.

26.
OSHA REGULATIONS: The Seller shall abide by Federal Occupational Safety and Health Administration (OSHA) regulations and the State of New Mexico Environmental Improvement Board occupational health and safety regulations that apply to the work performed under this Request. The Seller shall defend, indemnify and hold the University free and harmless against any and all claims, loss, liability and expense resulting from any alleged violation(s) of said regulation(s) including but not limited to, fines or penalties, judgments, court costs and attorneys fees.  
27.
PAYMENTS FOR PURCHASING: No warrant, check or other negotiable instrument shall be issued in payment for any purchase of services, construction or items of tangible personal property unless the central purchasing office or the using agency certifies that the services, construction or items of tangible personal property have been received and meet specifications.

28.
PROPOSAL NEGOTIATION: Offerors submitting proposals may be afforded an opportunity for discussion and revision of proposals. Revisions may be permitted after opening and prior to award for the purpose of obtaining best and final offer. Negotiations may be conducted with responsible offerors who submit offers found to be reasonably likely to be selected for award. The University is under no obligation to conduct discussions with any or all offerors.

29.
REFERENCES REQUIRED: Offeror must furnish a minimum of three (3) references of organizations and/or businesses that offeror has provided with similar services or items, unless otherwise stated in this Request for Proposal. The organization’s name, a contact, telephone number and a brief description of the items and services provided are required for each reference. Failure to submit the information may result in your offer being considered as non-responsive. Offeror, by furnishing these references, agrees to allow the University of New Mexico to contact any persons and/or organizations listed, and to utilize information obtained in evaluation of offer. The University of New Mexico reserves the right to check any references known to them and to use information received from such references in making the award.

30.
RELATIONSHIP OF PARTIES: The parties and their respective employees are at all times acting as independent contractors.  Contractor will not be considered an employee of the University of New Mexico for any purpose, including, but not limited to, workers’ compensation, insurance, bonding or any other benefits afforded to employees of the University.  Neither party has any express or implied authority to assume or create any obligation or responsibility on behalf of or in the name of the other party. 

31.
RELEASE UNM REGENTS: The Contractor shall, upon final payment of the amount due under the contract release Regents of the University of New Mexico, UNM Hospitals, their officers and employees and the State of New Mexico from liabilities, claims and obligations whatsoever arising from the contract.  The Contractor agrees not to purport to bind the University of New Mexico, UNM Hospitals or the State of New Mexico to any obligation not assumed in the contract by the Regents of University of New Mexico, UNM Hospitals or the State of New Mexico, unless the Contractor has express, written authority to do so, and then only within the strict limits of that authority.

32.
REMOVAL OF CONTRACTOR’S EMPLOYEE (S) The University of New Mexico may request that Contractor’s employee(s) be removed from the work under the contract for cause.  The University may immediately terminate, with written notice to Contractor, the services of any Contactor employee, if the University management believes in good faith that Contractor’s employee is unable to perform the services with reasonable skill.  Contractor’s agreement may also be terminated if Contractor’s liability insurance coverage is modified or terminated. 

33.
RETENTION OF RECORDS: Contractor will maintain detailed records indicating the date, time and nature of services provided under this Agreement for a period of at least five years after termination of this Agreement, and will allow access for inspection by The University of New, the Secretary for Health and Human Services, the Comptroller General and the Inspector General to such records for the purpose of verifying costs associated with provision of services under this Agreement.

34.
RIGHT TO PROTEST: Any offeror who is aggrieved in connection with this Request for Proposal may protest the award to the UNM Hospitals’ Purchasing Department as provided by the New Mexico State Procurement Code 13-1-172. The protest shall be submitted in writing within (15) fifteen calendar days after knowledge of the facts or occurrence-giving rise thereto.  

35.
RIGHT TO WAIVE MINOR IRREGULARITIES: The Evaluation Committee reserves the right to waive minor irregularities. The Evaluation Committee also reserves the right to waive mandatory requirements provided that all of the otherwise responsive proposals failed to meet the same mandatory requirements and the failure to do so does not otherwise materially affect the procurement. This right is at the sole discretion of the Evaluation Committee.

36.
SCHEDULE DELAYS: If after the award, the Seller becomes aware of possible problems that could result in delay in completion of the work on the agreed-to schedule, the Seller must immediately notify the Buyer or the designated representative. The initial notification of the delay may be verbal with a written confirmation, giving the probable cause and effect, with recommendations for alternate action. Nothing in this paragraph will be interpreted as relieving the Seller of its contractual obligations: however, failure to notify the University promptly will be a basis for determining the Seller responsibility in an otherwise excusable delay.

37.
SELLER GUARANTEE: The Seller shall guarantee all materials, equipment and workmanship furnished and/or installed under this Request to be free of defects and shall agree to replace solely at his expense, any and all defective equipment, parts, etc. within a one year period after the date of acceptance of the items and/or installation by the University. unless otherwise agreed to in writing at time of the award.

38.
SHORT LIST:  As stated in rule GSA 89-601 of the State of New Mexico Procurement Code regulations, for the purpose of evaluation and conducting discussion, a “short List” may be generated.  Proposals shall be ranked and the offerors who submitted the highest ranked proposals will be selected for the short list and will be eligible for evaluation and final negotiation.

39
STATE AND LOCAL ORDINANCES: The vendor shall perform work under this contract in strict accordance with the latest adopted version of all State and local codes, ordinances, and regulations governing the work involved.  
40.
TERMINATION PROVISION: This contract may be terminated by either party upon written notice delivered to the other party at least thirty days prior to the intended date of termination.  By such termination, neither party may nullify obligations nor liabilities already incurred for performance or for failure to perform prior to the date of termination.

41.
THIRD PARTIES: Nothing in this Agreement, express or implied, is intended to confer any rights, remedies, claims, or interests upon a person not a party to this Agreement.
42.
WAIVER: The contract shall contain a provision that states that no waiver of any breach of the Contract or any terms or conditions thereof shall be held to be a waiver of any other or subsequent breach; nor shall any waiver be valid, alleged or binding unless the same shall be in writing and signed by the party alleged to have granted the waiver.

43.
WARRANTY: Please state the warranty for equipment to be supplied under this Request. A copy of the warranty should be included in your submission.

SECTION II

THE UNIVERSITY OF NEW MEXICO STANDARD TERMS AND CONDITIONS

FOR GOODS AND SERVICES (Purchase Orders)

1. 
INSPECTION.   The University of New Mexico may inspect, at any reasonable time, any part of Seller’s plant or place of business, which is related to performance of this Purchase Order. Final inspection will be made at the destination upon completion of delivery of goods and services. Acceptance of delivery shall not be considered acceptance of the goods and/or services furnished. Final inspection shall include any testing or inspection procedures required by the specifications.

2. 
WARRANTIES.  Seller warrants the goods and/or services furnished to be exactly as specified in this Purchase Order, free from defects in Seller’s design, labor, materials and manufacture, and to be in compliance with any drawings or specifications incorporated herein and with any samples furnished by Seller.   All applicable UCC warranties express and implied are incorporated herein.

3.  
ELIGIBILITY FOR PARTICIPATION IN GOVERNMENT PROGRAMS.   Each party represents that neither it nor any of its management or any other employees or independent contractors who will have any involvement in the services or products supplied under this Agreement, have been excluded from participation in any government healthcare program, debarred from or under any other federal program (including but not limited to debarment under the Generic Drug Enforcement Act), or convicted of any offense defined in 42 U.S.C. Section l320a-7, and that each party, its employees and independent contractors are not otherwise ineligible for participation in federal healthcare programs. Further, each party represents that it is not aware of any such pending action(s) (including criminal actions) against each party or its employees or independent contractors. Each party shall notify the other immediately upon becoming aware of any pending or final action in any of these areas.

4. 
HEALTH INSURANCE PORTABILITY AND ACCOUNTABILITY ACT (HIPAA).  The parties agree to enter into a mutually acceptable amendment to this Agreement as necessary to comply with applicable federal laws and regulations governing the use and/or disclosure of individually identifiable health information.  Such amendment shall be entered into on or before the date by which hospitals are required to be in compliance with the privacy regulations promulgated under the Health Insurance Portability and Accountability Act of 1996.

5.
ACCEPTANCE AND REJECTION.   If prior to final acceptance, any goods or services are found to be defective or not as specified, or if the University of New Mexico is entitled to revoke acceptance of them the University of New Mexico may reject or revoke acceptance, require Seller to correct without charge within a reasonable time, or require delivery at an equitable reduction in price, at the University’s option. Seller shall reimburse the University of New Mexico for all incidental and consequential costs related to unaccepted goods or service. Notwithstanding final acceptance and payment, Seller shall be liable for latent defects, fraud, or such gross mistakes as amount to fraud. Acceptance of goods or services shall not waive the right to claim damages for breach of contract.

6. 
PAYMENT TERMS.  Upon written request from Seller for payment, The University of New Mexico shall, within 30 days, issue a written certification of complete or partial acceptance or rejection, with payment to follow within 30 days after certification of acceptance. Late payment charges shall be of 1 and 1/2 % per month.

7.
ASSIGNMENT.   This Purchase Order is assignable by the University.  Except as to any payment due hereunder, this Purchase Order is not assignable by Seller without written approval from the University.  

8.
CHANGES.   The University of New Mexico may make changes within the general scope of this Purchase Order by giving notice to Seller and subsequently confirming such changes in writing. If such changes affect the cost of, or the time required for performance of this Purchase Order, an appropriate equitable adjustment shall be made. No change by Seller shall be recognized without written approval of the University. Any claim of Seller for an adjustment under this Paragraph must be made in writing within thirty (30) days from the date of receipt by Seller of notification of such charge. Nothing in this Paragraph shall excuse Seller from proceeding with the performance of the Purchase Order as changed hereunder.

9. 
TERMINATION AND DELAYS.  The University of New Mexico may by written notice stating the extent and effective date, terminate this Purchase Order for convenience in whole or in part, at any time. The University of New Mexico shall pay Seller as full compensation for performance until such termination: (1) the unit or pro rata order price for the delivered and accepted portion: and (2) incidental damages, not otherwise recoverable from other sources by Seller, as approved by the University, with respect to the undelivered or unaccepted portion of this Purchase Order provided compensation hereunder shall in no event exceed the total Purchase Order price. Such amount will be limited to Seller’s actual cost, and may not include anticipated profits. The University of New Mexico shall not be liable for consequential damages. The University of New Mexico may by written notice terminate this Purchase Order in whole or in part for Seller’s default if Seller refuses or fails to comply with the provisions of this Purchase Order or fails to make progress so as to endanger performance and does not cure such failure within a reasonable period of time. In such event, the University of New Mexico may otherwise secure the materials, supplies or services ordered, and Seller shall be liable for damages suffered by the University of New Mexico thereby, including incidental and consequential damages. If after notice of termination, the University of New Mexico determines Seller was not in default, or if Seller’s default is due to failure of the University, termination shall be deemed for the convenience of the University. The rights and remedies of the University of New Mexico provided in this paragraph shall not be exclusive and are in addition to any other rights and remedies provided by law or under this Purchase Order as used in this paragraph, the word “Seller’ includes Seller and his sub-suppliers at any tier.

10. 
EQUAL OPPORTUNITY AND AFFIRMATIVE ACTION.  In performing the services required under this Purchase Order, each party shall be an equal opportunity employer and shall conform to all affirmative action and other applicable requirements; accordingly, each party shall neither discriminate nor permit discrimination in its operations or employment practices against any person or group of persons on the basis of race, age, religion, color, national origin, ancestry, sex, physical or mental handicap or medical condition, sexual preference, prior military involvement or in any other manner prohibited by law.

11. 
INDEMNIFICATION AND INSURANCE.   Seller assumes the entire responsibility and liability for losses, expenses, damages, demands and claims in connection with or arising out of any actual or alleged personal injury (including death) and/or damage or destruction to property sustained or alleged to have been sustained in connection with or arising out of the goods delivered by Seller or the performance of the work by Seller its agents, employees, subcontractors or consultants, except to the extent of liability arising out of the negligent performance of the work by or willful misconduct of the University. Seller shall indemnify, defend and hold harmless the University, its officers, agents, and employees from any and all liability for such losses, expenses, damages, demands, and claims and shall defend any suit or action brought against any or all of them based on any actual or alleged personal injury or damages and shall pay any damage costs and expenses including attorneys’ fees, in connection with or resulting from such suit or action.  Seller will also indemnify, defend and hold harmless the University of New Mexico against any joint and several liabilities imposed against the University of New Mexico with respect to strict product liability claims attributable to the fault of the Seller.

Seller agrees that it and its subcontractors will maintain general liability, product liability and property damage insurance in reasonable amounts (at least equal to the New Mexico Tort Claims Act limits) covering the above obligation and will maintain workers’ compensation coverage covering all employees performing under this Purchase Order on premises occupied by or under the control of the University.  The liability of the University of New Mexico will be subject in all cases to the immunities and limitations of the New Mexico Tort Claims Act, Sections 41-4-1 et seq. NMSA 1978, as amended.

12. 
INSTRUMENTALITIES.   Seller shall supply all equipment, tools, materials and supplies to accomplish the designated tasks except as set forth in this Purchase Order or its attachments.

13. 
PAYROLL OR EMPLOYMENT TAXES.   No federal, state, or local income, payroll or employment taxes of any kind shall be withheld or paid by the University of New Mexico with respect to payments to Seller or on behalf of Seller its agents or employees. Seller shall withhold and pay any such taxes on behalf of its employees as required by law. The payroll or employment taxes that are the subject to this paragraph include but are not limited to FICA, FUTA, federal personal income tax, state personal income tax, state disability insurance tax, and state unemployment insurance tax.  If Seller is not a corporation, Seller further understands that Seller may be liable for self-employment (social security) tax, to be paid by Seller according to law.

14. 
WORKERS COMPENSATION.  No workers compensation insurance has been or will be obtained by UNM on account of Seller or its employees or agents. Seller shall comply with the workers compensation laws with respect to Seller and Seller’s employees and agents.

15.
INDEPENDENT BUSINESS.  Neither Seller nor any of its agents shall be treated as an employee of the University of New Mexico for any purpose whatsoever.  Seller declares that Seller is engaged in an independent business and has complied with all federal, state and local laws regarding business permits and licenses of any kind that may be required to carry out the said business and the tasks to be performed under this Purchase Order. Seller further declares that it is engaged in the same or similar activities for other clients and that the University of New Mexico is not Seller’s sole or only client or customer.

16.  
SELLERS EMPLOYEES AND AGENTS.   Seller shall have complete charge and responsibility for persons employed by Seller and engaged in the performance of the specified work. The Seller, its agents and employees state that they are independent contractors and not employees of the University. Seller, its agents and employees shall not accrue leave, retirement, insurance, bonding or any other benefit afforded to employees of the University of New Mexico as a result of this Purchase Order.

17.
PATENT AND COPYRIGHT INDEMNITY.   Seller shall indemnify, defend and hold harmless the University of New Mexico against all losses, liabilities, lawsuits, claims, expenses (including attorneys’ fees), costs, and judgments incurred through third party claims of infringement of any copyright, patent, trademark or other intellectual property rights.

18.
DISCOUNTS.  If prompt payment discounts apply to this Purchase Order any discount time will not begin until the materials, supplies, or services have been received and accepted and a correct invoice received by the appropriate University Department. In the event testing is required prior to acceptance; the discount time shall begin upon the completion of the tests.

19.
PENALTIES.  The Procurement Code. Section 13-1-28 at seq. NMSA 1978 as amended imposes civil and criminal penalties for its violation. In addition, the New Mexico criminal statutes impose criminal penalties for bribes, gratuities and kickbacks.

20. TITLE AND DELIVERY.  Title to the materials and supplies passed hereunder shall pass to the University of New Mexico upon acceptance at the FOB point specified, subject to the right of the University of New Mexico to reject.  For any exception to the delivery date specified, Seller shall give prior notification and obtain approval thereto from the University’s Purchasing Department. Time is of the essence and the Purchase Order is subject to termination for failure to deliver on time. 

21. GOVERNING LAW.  This Agreement shall be construed in accordance with the laws of the State of New Mexico as they pertain to agreements executed and fully to be performed with New Mexico, or federal law where applicable, but in either case excluding that body of law relating to choice of law. 
22. OTHER APPLICABLE LAWS.   Any provision required to be included in an agreement of this type by any applicable and valid executive order, federal, state or local law, ordinance, rule or regulation shall be deemed to be incorporated herein.

23. CONFLICT OF INTEREST.  Seller shall disclose to the University of New Mexico Purchasing Department the name(s) of any University employee or member of the Board of Regents who has a direct or indirect financial interest in the Seller or in the proposed transaction.  A University employee (or Regent) has a direct or indirect financial interest in the Seller or in the proposed transaction if presently or in the preceding twelve (12) months the employee/Regent or a close relative has an ownership interest in the Seller (other than as owner of less than 1% of the stock of a publicly traded corporation); works for the Seller, is a partner, officer, director, trustee or consultant to the Seller, has received grant, travel, honoraria or other similar support from the Seller, or has a right to receive royalties from the Seller.  Seller shall file a Conflict of Interest Disclosure form with the UNM Hospitals Purchasing Department.  
24. AGREEMENT.  This Purchase Order is the sole and entire Agreement between the parties; any documents incorporated into this Agreement are listed explicitly on the front side of this Purchase Order, or are incorporated by Implication by the terms of this Purchase order.  Any terms inconsistent with or in addition to this Purchase order proposed by Seller are deemed rejected unless agreed to in writing by an appropriate University official.

25. FOB.  Unless stated otherwise, the price for goods is F.O.B. the place of destination, and the place of destination is the University’s designated campus address.  

26. VENDOR COMPLIANCE.    The Deficit Reduction Act became effective January 1, 2007, please review and understand the compliance training material under the “Vendor Compliance” section of the UNM Purchasing Department web page at www.unm.edu/~Purch

SECTION III

INSTRUCTIONS TO OFFERORS

1.
SEALED PROPOSALS: All proposals must be sealed and submitted in the prescribed manner and must be signed.  The authority of the individual signing the offer must be stated in the proposal.

2.
CONSIDERATION FOR ACCEPTANCE.  These discussions may relate to negotiated changes to any portion of the proposal.  However, proposals may be accepted without discussions.

3.
ACKNOWLEDGMENT OF ADDENDUM:  Offerors shall acknowledge receipt of any addendum to this Request for Proposal by identifying the addendum number and date in the space provided on the response form.

4.
PROPER ADDRESS: Indicate address to which UNM Hospitals purchase orders should be sent, if different from address shown on page 2 of this proposal.

5.
NUMBER OF COPIES OF PROPOSAL: Please submit the required number of copies of your proposal and all supporting documents stated on the cover sheet.  Failure to submit the required number of copies may result in your proposal being considered non-responsive.

6.
FAILURE TO RESPOND: UNM Hospitals reserves the right to remove from its Bidders’ list, the names of vendors who do not respond to Requests for Proposals. Vendors who are unable to respond to a specific Request for Proposal should notify the Buyer by letter if they wish to receive future Request for Proposals.

7.
CANCELLATION:  UNM Hospitals reserves the right to cancel without penalty, this Request, the resultant contract or any portion thereof for unsatisfactory performance or unavailability of funds.

8.
EMAIL ADDRESS OR TELEPHONE PHONE NUMBER FOR QUOTE CLARIFICATION: The Offeror must include an email address and a local or toll-free number for offer clarification.  Failure to do so may result in the quote being classified as non-responsive.

9.
CLARIFICATIONS:  Buyer shown on the cover sheet of this Request for Proposal shall be contacted concerning any clarification of instructions, terms and conditions, insurance, bonds, or offer preparation.  Technical clarifications should be addressed to the individual identified on the cover sheet.  Clarifications will be provided to all Offerors, in writing, as an addendum, and will be considered as part of this Request for Proposal.

10.
LATE SUBMISSIONS: Late submissions of proposals will be not be considered unless it is determined by  UNM Hospitals that late receipt was due solely to mishandling by  UNM Hospitals, after receipt by  UNM Hospitals, or whenever the proposal is the only offer received.  All other late submissions will be returned unopened.

11.
MODIFICATIONS:  Only Modifications received prior to time specified for the proposal opening will be accepted.  No modifications will be accepted following the opening.  Technical clarifications of the offer may be requested by the Buyer following the opening.  A late modification of an otherwise successful offer that makes its terms more favorable to the UNM Hospitals will be considered at any time it is received.

12.
PUBLIC INFORMATION: All information, except that classified as confidential, will become public information at the time that the contract is awarded.  Confidential information must be marked “CONFIDENTIAL” in red letters in the upper right hand corner of the sheets containing the confidential information.  Price and information concerning the specifications cannot be considered confidential.

13.
PAYMENT DISCOUNTS:  UNM Hospitals will take advantage of payment discounts offered whenever possible; however, payment discounts will not be used as a means to determine the lowest responsible offeror.

14.
PERIOD FOR OFFER ACCEPTANCE: Except as otherwise noted in this RFP, the offeror agrees, if its offer is accepted within 90 calendar days of the final closing date, to furnish any and all item(s) and/or services at the prices set forth in its offer, delivered to the designated point(s) within the specified time in the delivery schedule.  Failure to comply may result in removal from UNM Hospitals’ Offeror’s List.  UNM Hospitals reserves the right to request an extension of this time period.

15.
REJECTION OF OFFERS:  UNM Hospitals reserves the right to award offers based on evaluation criteria contained herein, to reject any and all offers or any part thereof, and to accept the offer that is in the best interests of UNM Hospitals.

16.
TAXES:   UNM Hospitals is exempt from Federal Excise Taxes and from New Mexico Sales Taxes on materials, except construction materials used by a contractor.  Services are not exempt.  Taxes on services should be included as a separate line item and not included in any base price quote.  Applicable taxes are excluded from the proposal evaluation.  A non-taxable transaction certificate will be provided upon request.

17.
WITHDRAWAL OF OFFERS:  Offers may be withdrawn by written notice, Fax, telegram or in person, by an authorized representative of the offeror, any time prior to the award.

18. CASH DISCOUNTS:  UNM Hospitals is a non-profit, state supported educational institution and available discounts should be noted in offeror’s response.

19. FAX OFFERS:  Fax offers will not be accepted unless an authorized representative of the Offeror is present to receive and seal the quote.  UNM Hospitals' personnel will not accept fax offers. 
20.
SUBMISSIONS OF DRAWINGS/LITERATURE: The submission of drawings and literature to be used in the evaluation of the offer must be made by the closing date and time to be considered.  All submissions shall be made at no expense to the UNM Hospitals.  Returns of submissions will be made only when requested by the Offeror and only at Offeror’s expense.
21.
RFP COSTS: The costs of preparing responses to this RFP, and the costs associated with preparing for and participating in the demonstrations is to be borne by the offerors.  The costs that result from participating in this procurement may not be charged, directly or indirectly, against any contract signed with UNM Hospitals. UNM Hospitals Employees are Final Judges

SECTION IV

INSURANCE REQUIREMENTS
CERTIFICATES OF INSURANCE:

The Contractor shall furnish the Owner one copy each of Certificates of Insurance herein required for each copy of the Agreement, showing coverage, limits of liability, covered operations, effective dates of expiration of policies of Insurance carried by the Contractor. The Contractor shall furnish to the Owner copies of limits. The Certificate of insurance shall be in the form of Alas Document G-705 or similar format acceptable to the Owner. Such certificates shall be filed with the Owner and shall also contain the following statements:

1.
‘The Regents of The University of New Mexico, The University of New Mexico, its agents, servants and employees are held as additional insured.”

2.
“The insurance coverage certified herein shall not be canceled or materially changed except after forty five (45) days written notice has been provided to the owner.”

COMPENSATION INSURANCE:

The Contractor shall procure and shall maintain during the life of this contract Worker’s Compensation Insurance as required by applicable State law for all Contractor’s employees to be engaged at the site of the project under this project and in case of any such work sublet, the Contractor shall require the subcontractor or sub subcontractor similarly to provide Worker’s Compensation Insurance for all the subcontractor’s or sub subcontractor’s Workers which are covered under the Contractor’s Worker’s Compensation insurance. In case any class of employee engaged in work on the project under this contract is not protected under a Worker’s Compensation Statute, the Contractor shall provide and shall cause each subcontractor or sub subcontractor to provide Employer’s insurance in an amount of not less than $500,000.

CONTRACTOR’S PUBLIC LIABILITY INSURANCE:

The contractor shall maintain liability insurance coverage “equal to the maximum liability amounts set forth in the New Mexico Tort Claims Act Section 41-4-1 et. seq. NMSA 1978”. The insurance must remain in force for the life of the contract Including all contract extensions or renewals. The limits effective July 1, 1992 are:

$400,000 per person/$750,000 per occurrence plus $300,000 for medical and $100,000 for property damage for a total maximum liability of $1,150,000 per occurrence.

CONTRACTOR’S VEHICLE LIABILITY INSURANCE:

The Contractor shall procure and shall maintain during the life of this contract Vehicle Liability Insurance coverage “equal to the maximum liability amounts set forth in the New Mexico Tort Claims Act Section 41-4-1 et. Seq.  NMSA 1978”. The insurance must remain in force for the life of the contract including all contract extensions and renewals. The limits effective July 1, 1992 are:  Bodily injury $750,000 Each Occurrence--Property Damage $100,000 Each Occurrence
CONTRACTOR’S AND SUB SUBCONTRACTOR’S PUBLIC AND VEHICLE LIABILITY INSURANCE:
The Contractor shall either:
1.
Require each subcontractor or sub subcontractor to procure and maintain during the life of the subcontract or sub subcontract public Liability Insurance of the types and amounts specified above or,

2.
Insure the activities of the subcontractors of sub subcontractors in the Contractor’s Policy as required under this Article.

GENERAL
All insurance policies are to be issued by companies authorized to do business under the laws of the state in which work is to be done and acceptable to owner.

The Contractor shall not violate, permit to be violated, any conditions of any said policies, and shall at all times satisfy the requirements for the insurance companies writing said policies.

SECTION  IV
THE UNIVERSITY OF NEW MEXICO SUPPLIER CONFLICT OF INTEREST AND

DEBARMENT/SUSPENSION CERTIFICATION FORM

THE FOLLOWING MUST BE CERTIFIED IF THIS PURCHASE ORDER IS $10,000 OR GREATER

CONFLICT OF INTEREST

The authorized Person, Firm and/or Corporation states that to the best of his/her belief and knowledge: No employee or Regent of The University of New Mexico (or close relative), with the exception of the person(s) identified below, has a direct or indirect financial interest in the Vendor or in the proposed transaction.  Vendor neither employs, nor is negotiating to employ, any University of New Mexico employee, Regent or close relative, with the exception of the person(s) identified below.  Vendor did not participate, directly or indirectly, in the preparation of specifications upon which the quote or offer is made.  If the Vendor is a New Mexico State Legislator or if a New Mexico State Legislator holds a controlling interest in Vendor, please identify the legislator: _____________________________________ List below the name(s) of any University or New Mexico employee, Regent or close relative who now or within the preceding 12 months (1) works for the Vendor; (2) has an ownership interest in the Vendor (other than as an owner of less than 1% of Vendor’s stock, if Vendor is a publicly traded corporation); (3) is a partner, officer, director, trustee or consultant to the Vendor; (4) has received grant, travel, honoraria or other similar support from Vendor; or (5) has a right to receive royalties from the vendor.____________________________________________________________________________________

DEBARMENT/SUSPENSION STATUS

The Vendor certifies that it is not suspended, debarred or ineligible from entering into contracts with the Executive Branch of the Federal Government, or in receipt of a notice or proposed debarment from any Agency.  The vendor agrees to provide immediate notice to The UNM Hospital Purchasing Department in the event of being suspended, debarred or declared ineligible by any department or federal agency, or upon receipt of a notice of proposed debarment that is received after the submission of the quote or offer but prior to the award of the purchase order or contract.
CERTIFICATION

The undersigned hereby certifies that he/she has read the above CONFLICT OF INTEREST and DEBARMENT/SUSPENSION Status requirements and that he/she understands and will comply with these requirements.  The undersigned further certifies that they have the authority to certify compliance for the vendor named and that the information contained in this document is true and accurate to the best of their knowledge.  

Signature: ____________________________Title: _____________________________Date: __________________​​​



Name Typed: __________________________________
Company Name_____________________________________

Address
______________________________________
City/State/zip: ______________________________________

THE FOLLOWING MUST BE CERTIFIED IF THIS PURCHASE ORDER IS $100,000 OR GREATER:

CERTIFICATION AND DISCLOSURE REGARDING PAYMENTS TO INFLUENCE CERTAIN FEDERAL TRANSACTIONS (APRIL 1991)

A. The definitions and prohibitions contained in the clause, at FAR 52.203-12, Limitation on Payments to influence Certain Federal Transactions, I included in this solicitation, are hereby incorporated by reference in paragraph (b) of this certification.

B. The offeror, by signing its offer, hereby certifies to the best of his or her knowledge and belief that on or after; December 23, 1989

1) Federal appropriated funds have not been paid and will not be paid to any person for influencing or attempting to Influence an officer or employee of any agency, a Member of Congress, an officer or employee of Congress, or an employee of a Member of Congress on his or her behalf in connection with the awarding of any Federal contract, the making of any Federal grant, the making of any Federal loan, the entering into any cooperative agreement, and the extension, continuation, renewal, amendment or modification of any Federal contract, grant, loan, or cooperative agreement; 

2) If any funds other than Federal appropriated funds (including profit or fee received under a covered Federal Transaction) have been paid, or will be paid, to any person for influencing or attempting to influence an officer or employee of any agency, a Member of Congress, and officer or employee of Congress, or an employee of a Member of Congress on his or her behalf in connection with this solicitation, the offeror shall complete and submit, with its offer, OMB standard form LLL, Disclosure of Lobbying Activities, to the Contracting Officer; and He or she will include the language of this certification in all subcontract awards at any tier and require that all recipients of subcontract awards in excess of $100,000 shall certify and disclose accordingly.

3) (c) Submission of this certification and disclosure is a prerequisite for making or entering into this contract imposed by section 1352, title 31, United States Code.  Any person who makes expenditure prohibited under this provision or who fails to file or amend the disclosure form to be filed or amended by this provision shall be subject to a civil penalty of not less than $10,000 and not more than $100,000 for each such failure.

CLEAN AIR ACT AND FEDERAL WATER POLLUTION CONTROL ACT
The undersigned company agrees to comply with all applicable standards, orders or regulations issued pursuant to the Clean Air Act (42 U.S.C. 7401 et seq.) and the Federal Water Pollution Control Act as amended (33 U.S.C. 1251 et seq.)
CERTIFICATION
The undersigned hereby certifies that he/she has read the above CERTIFICATION AND DISCLOSURE REGARDING PAYMENTS TO INFLUENCE CERTAIN FEDERAL TRANSACTION (APR 1991) and CLEAN AIR ACT AND FEDERAL WATER POLLUTION CONTROL ACT requirements and that he/she understands and will comply with these requirements.  The undersigned further certifies that they have the authority to certify compliance for the vendor named below.
Signature: __________________________________________Title: ___________________________________ Date: _________

Name Typed: _______________________________________Company: _____________________________________________

Address: __________________________________________City/State/zip: __________________________________________

SECTION VI
SCOPE OF WORK/SPECIFICATIONS

Purpose: The purpose of the Interventional Radiology (IR) RFP is to identify and select the optimal solution for the growing service in Radiology, A partnership for future development. The goal is to select a product line which we can anticipate upgrading all 4 labs over the next 4 years. Our initial intent this fiscal year is to select and acquire a bi-plane Interventional room solution. 
I. Description of Optimal System
a. The Selected vendor shall provide the highest quality systems to perform Adult and pediatric procedures. The system is to be of a high performance design to satisfy the most demanding medical and technological requirements in all as ranges of interventional radiology to include neuroradiology (Cerebral and Vertebral anatomy).  
b. The response to this RFP shall include specifications and pricing for a Bi-plane interventional solution. 

c. The system must provide a digital acquisition process which can function and be reviewed independently of each room. 

d. The system will have a table mount MEDRAD power injector included. 

e. The solution will provide an ultrasound system for various applications in the IR realm. 

f. The system must include DICOM Send, DICOM Print, DICOM store, image export, and modality worklist. Vendor to provide a DICOM compliance statement. In particular it must support the current and future objects for patient dose.  
g. The systems shall provide high patient volume capabilities that can be achieved with minimum personnel requirements.

h. Each system must feature a complete line of measures to considerably reduce the radiation exposure for the patient population as well as staff. 

i. Each system shall produce high resolution; fast frame rates and high tube load so that even the tiniest of vessels are clearly visible in high contrast with the smallest amount of radiation. 

j. Image data transmission shall be industry standard, DICOM compliant for connection to the hospital network which includes the need to interface with the existing CERNER information system. 

k. Considering the overall plan to improve and grow over the next 4 years, provide a solution to ensure models stay on most current platform, provide an upgrade solution.

l. Vendor to understand what we currently have and provide solutions to optimize, i.e.…upgrade, trade-in, etc…

II. Minimal accepted equipment specifications
This section describes basic requirements for the room.  The requirements are a brief, not necessarily all-inclusive, outline of equipment and features to be provided by the successful VENDOR.  It is a guideline rather than an itemization.  The provided response shall include all equipment, accessories, and features necessary for performing the scope of services required by the UNMH IR Department.

It is realized that each Vendor’s system may differ from these requirements.  It is required, however, that whenever a variance from these requirements occurs, that the item(s) proposed shall meet or exceed the specified level of performance or characteristics.  

If a specified requirement is unavailable, alternative(s) must be proposed; the VENDOR is also strongly encouraged to comment on its future availability by providing details.
a. Table
i. The table will be digital 
ii. The table shall have motorized or float top movement, direct patient access from all sides.
iii. The table will provide support to a weight limit of 400 lbs. up to 650 lbs.
iv. Specify height adjustment mechanism
v. Specify table size
vi. Identify table rotation/swivel for peripheral vascular cases
vii. Specify Cardio-Pulmonary Resuscitation weight capacity
viii. Identify patient mattress
1. table top mattress

2.  mattress and/or sponges for cerebral angiography
ix. Vendor to quote all patient support accessory options available separately.
x. Specify table top transmission.  Optimally it should not exceed 0.6 mm Al equivalent.

xi. Specify upper/lower body radiation protection for staff- both right and left side operation.
xii. Identify industry standard examination lamp 

xiii. Describe runoff or bolus chase capabilities
b. Controls
i. Require optimal solution for control modules in the lab for image review, specify functionality.
1. Table side controls, vendor to state options regarding protocol manipulation if available 
2. Mobile table controls

3. C-Arm angulations
a. Image intensifier zoom size
b. Collimator control
i. All collimator and filter controls shall be at tableside.

ii. Specify collimators, i.e.… for angiography with square blade and wedge-shaped diaphragm plate
iii. Describe functions of collimators and capabilities during fluoro off to reduce exposure

iv. Filter functions shall include:
1. Graduated straight edge or contour-type filter
2. Each filter should individually adjust and rotate 180 degrees.
3. Automatic synchronous rotation of the camera and collimator unit to compensate for image rotation at different examination positions of the support stand
4. Rotation override available tableside 
4. Filter control-Automatic
5. Position programming and recall
6. Reference image selection and recall
7. Collision protection 
8. Set frame rates
9. Fluoro 5 minute  timer reset
10. Dynamic comparison functionality

11. Multiple image display

12. Image annotation availability via letters, symbols and arrows

13. Post processing
14. Specify the standard and optional analysis which can be performed
15. Storage and display of dynamic fluoro sequences
16. Does system display patient dose information

ii. List all functions from control room 
iii. Identify Data display and ease of use

iv. Provide an intercom system for communication between exam and control room which can be increased as more rooms are acquired. 
c. Generator
i. Generator to provide high frequency design and reliable consistent power.
ii. Be suitable for every possible vascular examination. 
iii. Automatic calculation and optimization of fluoroscopic data.  
iv. Specify  kilowatt of generator output 
v. Specify the generator’s maximum kVp
vi. Automatic KV and mA control 
vii. Describe generator capable frame rate per second for both single plane and bi-plane
viii. Describe solution for varying or minimal radiation dosage for pediatric and long interventions
d. Tube
i. Specify tube anode heat storage capacity.
ii. List focal spot sizes 

e. Image Detectors 
i. Require Flat panel- Specify size
ii. Specify detector output matrix and dose efficiency
iii. Vendor to provide optimal solution for UNMH work load and case mix

iv. List imaging and zoom modes in centimeters
v. Describe low and high contrast resolution in each mode.  Specify how these were measured.

vi. The Image Detector shall be of sufficient cm sizing to allow the operator to match the field size to the relevant anatomy or smallest micro catheter and/or sheath for excellent image quality at the lowest possible dose.
vii. Must meet/exceed standard requirements for diagnostic and interventional. Neuro-Interventions are a must in this room. Please specify and include what will provide optimal solution. 

f. Imaging System
i. System must provide display of bi-plane live non-subtracted and live subtracted fluoro simultaneously
ii. Describe contrast enhancement and noise reduction in image without frame integration
iii. The last fluoroscopic image must be stored as an image and as a run

iv. List Real-time DSA speeds

v. Trace subtraction fluoroscopy must be included
vi.  Road mapping must be included. Describe road mapping capabilities applied to an existing DSA run. 

vii. 3-D image reconstruction must be available. Describe 3-D rotational angiography 
1. State the number of images acquired
2. Describe the rotational gantry speeds
3. Describe the anatomical coverage possible
4. Define the total acquisition time for maximum quality reconstructions, the reconstruction time, and what features are available.  

5. Describe where and how the reconstructions are performed 

viii. 3-D images must track together as the gantry position is varied.
ix. Describe reconstructed cross-sectional imaging obtained fluoroscopically and capabilities

x. Is the system capable of selecting an image from a prior run and using as an overlay 

g. C-Arm (s) Performance 
i. Multi-directional isocentric imaging-Capable of motorized and manual pivot movements so the operator can move the gantry quickly from the head of the patient or to either side of the patient. 

ii. Gantry must allow  three-sided patient approach

iii. List the C-Arm projection minimum/maximums for “RAO, LAO, Cranial, and Caudal”.
iv. Specify/distinguish lateral C-arm performance 

v. Specify rotational speed for LAO/RAO
vi. Specify rotational speed for Cranial/Caudal
vii. List number of pre-program positioning
viii. Describe focus-isocenter capabilities
ix. Single joystick for all patient-angle oriented c-arm movements with collision monitoring

h. Monitors/Monitor Suspension (prefer 8+ bank set up)
i. Monitors to be of high resolution-specify
ii. Exam room monitors shall be a minimum of 18 inches- specify
1. Needs: 2 live feed, 2 reference, 1 3D, 1 PACS images, 1 Telemetry, 1 hemodynamic

2. Must integrate to existing systems, PACS, telemetry, hemodynamic. Vendor to assess UNMH current products.

iii. Control room monitors shall be a minimum of 19 inches, minimum 2-specify size
iv. Monitors brightness shall automatically adjust to surrounding procedure room light levels and provided an anti-reflective coating to minimize glare.

v. Hemodynamic monitors and image monitors must be on the same bank of monitors.

vi. Ability to import PACS images for reference, please describe

vii. Patient monitoring to be visualized on bank of monitors. Vendor to evaluate IR dept. current monitors for integration or provide optimal solution.
viii. Option 2 Remote monitors to be located in a remote location
i. General
i. Vendor will specify room size required (Free from Support Structures), vendor to state minimum floor to ceiling requirements
ii. The vendor will provide typical installation times and condition of room needed to start installation.
iii. Vendor to specify room lighting requirements to optimize image quality. 

iv. Vendor shall supply UPS back up for all systems requiring a minimum of 30 minute run time to eliminate possible loss of raw and processed images. List all options and provide recommendations to meet the need of UNMH Interventional Radiology.

v. Vendor to specify and provide protective shielding for staff 

vi. Describe your current and future support for patient recording in the DICOM header.  As this is changing you must agree to upgrade the system to support the current version of this information.

vii. Describe your specification of patient dose and cumulative patient dose and provide vendor’s specific enhancements and options for minimum dose without impact to image quality. Vendor shall provide all options which integrate patient dose information in an electronic format which can be sent to a recommended data base. 
j. Accessories

i. Radiation shield on a track and a table mounted radiation shield to be provided-specify
ii. Ceiling mounted exam lamp to be included-specify
iii. Table mounted pressure injector (prefer mount on a table side-rail)-specify
iv. Ultrasound device to enhance entire interventional imaging capabilities (probes and software capable of vascular access and body imaging). Must be DICOM capable include modality worklist -specify
III. System Acceptance
a. Verification that the required equipment and services meet the RFP requirements should be accomplished through an Acceptance Test, final payment will not be released prior to acceptable acceptance test and warranty start date will commence on same date.
b. The vendor should identify all required materials and supplies for the acceptance test.
c. Vendor will suggest acceptance tests required to meet their specifications Include in the quotation.

IV. Training
a. Appropriate manuals or documentation should be provided for each delivered system component.
b. QC procedures required to maintain operation within the requirements of this RFP should be supplied for each component.
c. All manuals should be updated appropriately to reflect each new software release and implementation phase.
d. The vendor should include a comprehensive training plan that includes:
i. On-site vendor training staff
ii. Courses offered
iii. Summary of course content
iv. Frequency of course offerings
v. Training programs for: technologists and service personnel
vi. Continuous, computer based training materials and programs.
1. Vendor should indicate if training qualifies as “continuing education.”
V.  Warranty and Maintenance
a. Warranty coverage should be provided at no charge for one year from the date of system acceptance.
b. Warranty should cover hardware and software of all system components supplied by vendor.
c. Warranty should cover parts and labor.
d. System downtime should be defined as a period when images cannot be acquired from an acquisition system, sent to or received from a storage system, or viewed on a workstation.
e. Total system uptime should be greater than 98% calculated monthly, excepting preventive maintenance periods.
f. The vendor should describe the spare parts stocking strategy that will be employed during the warranty period.
g. The vendor should perform all the preventive maintenance for during the warranty period so as to minimize workflow disruption. The vendor will provide all preventative maintenance after standard business hours M-F 0700-1700.
h. The vendor will supply personnel support during the warranty period with plans to phase these responsibilities over to the maintenance agreement and local service.
i. The vendor should submit all available Service Level Agreements, to include software/hardware upgrade options associated with each service level.
j. The vendor should include in the proposal the specifics of a maintenance contract including the option for software/hardware refresh agreement.
k. The maintenance contract should include options for one, two, three, four or five year’s duration following warranty.
l. The maintenance contract should cover parts and labor.
m. The maintenance contract should cover hardware and software of each system component supplied by vendor.
n. The cost for each major system component should be listed separately in the maintenance contract proposal.
o. The Vendor should outline the methodology used for placing support calls for hardware and application service.
p. The vendor will specify training for in house x-ray service staff.  UNMH x-ray service is provided by outside contract.  Currently Philips Medical System provides this support.
q. The vendor will describe response time for service call. For example if a patient is on the table and the machine fails.  How would this be handled?

r. The vendor will describe the service organization, training of staff to be servicing UNMH.

s. Vendor to provide an upgrade solution to keep room(s) with same software and hardware versions. 
t. Vendor service personnel must participate in a UNMH environment of care permit process and policies regarding access into organization.

u. Selected vendor will be required to provide UNMH Clinical Engineering Director with service manual, operations manual, and parts manuals. This documentation can be provided in paper or electronic media. 
VI. Spare Parts
a. The vendor should provide a guarantee that the hospital will be able to purchase any required spare parts from the vendor for five years from the time the component has ceased being manufactured.
b. Upgrades/Patches that become available during the warrantee period will be installed at not cost including all parts and labor.
c. How often are upgrades/patches installed? Is this part of maintenance? Who applies these? Can these be done with no downtime?
VII. Installation
a. Must be completed and accepted by June 30, 2010 otherwise instructed by UNMH. UNMH will provide a 60-day notice of install date.
Questions: Address all questions via email to the individuals listed below no later than 4:00 PM,  Friday May 22, 2009
Buyer: 

Lynn Nelson

email: salnelson@salud.unm.edu
Technical: 

Gordon Weimer

email: gweimer@salud.unm.edu
SECTION VII.

PROPOSAL EVALUATION CRITERIA

GENERAL: This section describes the criteria to be used for analyzing and evaluating the various proposals.  Although cost will be a factor in proposal evaluation, it is specifically a consideration of secondary importance to the need for providing a high quality system and a professional technical supporting staff to provide warranty and maintenance services.

	Evaluation Factor
	 Factor Weighting

	· Response conformance to RFP
	10%

	· Functional Requirements compliance
	20%

	· Technical and Operational soundness of proposed configuration
	35%

	· Competitiveness of System Pricing, Terms and Conditions offered
	35%


EVALUATION CRITERIA: All responses to this Request for Proposals become the property of UNMH and will be held confidential, with the exception of the successful proposal(s), which will become public information upon completion of UNMH contract negotiation process. UNMH reserves the right to judge the presentation of the firms submitting proposals in the evaluation and selection of the successful proposal. 

ORAL PRESENTATION: 

RFP Respondents being considered for final selection may be afforded a formal presentation or requested to participate in a site visit.  Respondents should be prepared to answer questions concerning their proposal during this presentation.  UNM Hospital reserves the right to award without formal presentations.
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